Safety Committee Meeting Minutes				Date:  3.12.2025

Present:  Dave Hemp, Brad Besse

Injury Update:

Drills:  None to report

Inspections:  Science Lab inspections:

Completed by Mr. Hemp over semester break.

- One fume hood in ST 256 taken out of service – did not pass airflow test.

Action:  Notified maintenance and Dept. Chair.  Looking at options

- Eyewash/safety showers not being tested monthly.

Action:  Notified Lab Assistant to please test these monthly.

COPE Inspection:  Insurance underwriter inspection, I was unable to attend.  I provided information regarding security issues, Fire response and emergency response plans.

Dave’s report:  Light bulb recycling – 833 bulbs recycled, $736.62

NARCAN placed in two AED boxes – RA and SC

Round Robin:  Brad: Nothing to report